## Sheet1
Schedule of Prices
Proponents need to provide all inclusive unit prices for the proposed services as well as additional prices for any additional work. The spreadsheet will automatically update the unit prices based on any annual escalators provided below. Proponents need to enter their unit prices into the highlighted boxes below but may override any unit price box as required.
Site: Soames Creek – Low Flow Weir (Upstream of Marine Drive)
Component | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Comments
Annual Calibration Visit
OTT PLS pressure transducer (0–4 m range) – bench cal | 0 | 0 | 0 | 0 | 0 | Remove or isolate sensor; multi‑point wet/pressure calibration vs. traceable standard; adjust coefficients in logger if required; issue calibration certificate.
CR310 datalogger – analog input calibration/verification | 0 | 0 | 0 | 0 | 0 | Verify voltage/current input channels used by PLS against calibrated signal source; check scaling, units, time stamps, and logging program.
CR310 datalogger – firmware & program review | 0 | 0 | 0 | 0 | 0 | Confirm firmware up to date; back up and validate program; verify logging intervals, alarms, and data storage.
Power system inspection (if applicable: batteries, solar) | 0 | 0 | 0 | 0 | 0 | Check battery voltage, charge controller, fuses, and wiring; verify stable power to PLS, CR310, and modem.
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Quarterly verification & telemetry check visit (x4 per year) Prices to be provided will be for each visit
OTT PLS pressure transducer – in situ verification | 0 | 0 | 0 | 0 | 0 | Field check against manual staff gauge readings and/or portable reference sensor; document offset; apply correction in CR310 if needed.
Cellular modem – functional test | 0 | 0 | 0 | 0 | 0 | Verify modem power, signal strength, SIM status, and connection to network; confirm automatic reconnect behavior.
Telemetry link to Granthams WTP PLC – end‑to‑end test | 0 | 0 | 0 | 0 | 0 | Confirm data are received at WTP PLC; check Modbus mapping, registers, scaling, and units; verify no dropped or stale values.
Modbus communications check (CR310 ↔ modem/PLC) | 0 | 0 | 0 | 0 | 0 | Confirm Modbus address, baud rate, parity, and register map; test read/write where applicable; document configuration.
Time synchronization check (CR310 vs. WTP PLC/SCADA) | 0 | 0 | 0 | 0 | 0 | Compare logger clock to WTP PLC/SCADA time; adjust drift; confirm time zone and daylight‑saving settings.
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Ad-Hoc Call Out / Trouble Shooting
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Travel Costs: | 0 | 0 | 0 | 0 | 0
Site: Soames Creek – Chapman Creek – Reach 3
Component | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Comments
Annual Calibration Visit
OTT PLS vented pressure transducer (±2 mm accuracy) – bench calibration | 0 | 0 | 0 | 0 | 0 | Multi‑point pressure calibration against traceable standard; verify vent tube integrity; update calibration coefficients; issue certificate.
CR310 datalogger – analog input verification | 0 | 0 | 0 | 0 | 0 | Verify voltage/current input channels; confirm scaling, units, and data program; check time stamps and memory.
CR310 firmware & program audit | 0 | 0 | 0 | 0 | 0 | Confirm firmware version; back up program; validate logging intervals, alarms, and data storage.
Battery power system inspection (deep‑cycle battery) | 0 | 0 | 0 | 0 | 0 | Measure battery voltage under load; inspect terminals, wiring, fuses; check charge controller (if solar present).
Stilling pipe inspection (steel pipe anchored to boulder) | 0 | 0 | 0 | 0 | 0 | Inspect for blockage, sediment, biofouling, corrosion, or deformation; verify sensor mounting height; clean as required.
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Quarterly verification & telemetry check visit (x4 per year) Prices to be provided will be for each visit
OTT PLS – in situ verification | 0 | 0 | 0 | 0 | 0 | Compare sensor readings to staff gauge or reference water level; check vent tube desiccant; document offsets; apply corrections in CR310 if required.
Cellular telemetry modem – functional test | 0 | 0 | 0 | 0 | 0 | Verify power, antenna, signal strength, SIM status; confirm successful data transmission and reconnect behavior.
Telemetry link end‑to‑end test (station → server/SCADA) | 0 | 0 | 0 | 0 | 0 | Confirm data receipt at host system; check register mapping, scaling, and units; verify no stale or missing data.
Time synchronization check (CR310 vs. SCADA/server) | 0 | 0 | 0 | 0 | 0 | Compare and correct clock drift; verify time zone and DST settings.
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Ad-Hoc Call Out / Trouble Shooting
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Travel Costs: | 0 | 0 | 0 | 0 | 0
Site: Chapman Creek – WSC / Downstream of Intake
Component | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Comments
Annual Calibration Visit
Vega C21 Radar Level Transmitter – bench calibration | 0 | 0 | 0 | 0 | 0 | Verify radar accuracy using certified reference distances; check beam alignment; update configuration; issue calibration certificate.
Keller Level Pressure Transmitter – bench calibration | 0 | 0 | 0 | 0 | 0 | Multi‑point pressure calibration against traceable standard; verify zero/span; update scaling in PLC; issue certificate.
Stilling well inspection (suspended from bedrock‑mounted shed) | 0 | 0 | 0 | 0 | 0 | Inspect for sediment, blockage, corrosion, deformation; verify stilling well verticality; clean and flush as required.
PLC‑based data acquisition system – I/O calibration | 0 | 0 | 0 | 0 | 0 | Verify analog input channels with calibrated signal source; confirm scaling, units, alarms, and data logging logic.
PLC firmware, program, and communications audit | 0 | 0 | 0 | 0 | 0 | Review ladder/logic program; back up configuration; verify Modbus/telemetry mapping and data integrity.
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Quarterly verification & telemetry check visit (x4 per year) Prices to be provided will be for each visit
Vega C21 – in situ verification | 0 | 0 | 0 | 0 | 0 | Compare radar readings to manual staff gauge and benchmarked water level; confirm mounting stability; document offsets.
Keller Transmitter – in situ verification | 0 | 0 | 0 | 0 | 0 | Compare pressure‑derived water level to radar and manual readings; check vent tube (if vented); inspect cable and fittings.
Telemetry/SCADA end‑to‑end data verification | 0 | 0 | 0 | 0 | 0 | Confirm live data transmission to SCADA/server; check register mapping, scaling, timestamps, and data freshness.
Permanent benchmark verification | 0 | 0 | 0 | 0 | 0 | Survey check of benchmarks; confirm stability; verify manual gauge readings relative to benchmarks; document any movement.
Time synchronization check (PLC ↔ SCADA/server) | 0 | 0 | 0 | 0 | 0 | Compare and correct clock drift; verify time zone and DST settings.
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Ad-Hoc Call Out / Trouble Shooting
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Travel Costs: | 0 | 0 | 0 | 0 | 0
Site: Chapman Creek Upper
Component | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Comments
Annual Calibration Visit
OTT PLS vented pressure transducer – bench calibration | 0 | 0 | 0 | 0 | 0 | Multi‑point pressure calibration against traceable standard; verify vent tube integrity; update coefficients; issue calibration certificate.
CR310 datalogger – analog input verification | 0 | 0 | 0 | 0 | 0 | Verify voltage/current input channels; confirm scaling, units, logging intervals, and program logic; check timestamps and memory.
CR310 firmware & program audit | 0 | 0 | 0 | 0 | 0 | Confirm firmware version; back up program; validate alarms, data storage, and communication settings.
Telemetry-ready enclosure inspection | 0 | 0 | 0 | 0 | 0 | Inspect enclosure seals, cable glands, grounding, and mounting; verify environmental protection above high-water level; check internal wiring and strain relief.
Battery power system inspection (deep-cycle battery) | 0 | 0 | 0 | 0 | 0 | Measure battery voltage under load; inspect terminals, wiring, fuses; check charge controller (if solar present).
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Quarterly verification & telemetry check visit (x4 per year) Prices to be provided will be for each visit
OTT PLS – in situ verification | 0 | 0 | 0 | 0 | 0 | Compare sensor readings to manual gauge or reference level; inspect vent tube and desiccant; document offsets; apply corrections in CR310 if needed.
Telemetry system functional test | 0 | 0 | 0 | 0 | 0 | Confirm modem or telemetry hardware (if installed) is operational; verify signal strength, data transmission, and reconnect behavior.
Time synchronization check (CR310 ↔ server/SCADA) | 0 | 0 | 0 | 0 | 0 | Compare and correct clock drift; verify time zone and DST settings.
Documentation & reporting | 0 | 0 | 0 | 0 | 0
Ad-Hoc Call Out / Trouble Shooting
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Hourly Rates (Describe) : | 0 | 0 | 0 | 0 | 0
Travel Costs: | 0 | 0 | 0 | 0 | 0
Separate Prices
Component | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Comments
Hydrometric Data Collection Processing Conversion - Cost per Site | 0 | N/A | N/A | N/A | N/A | Two (2) of the sites have the data pass through a third party system first and we need it to pass through the Regional Districts system first as it is the system of record.
Maintenance, Upgrade and Repair Costs - Labour - Cost Per Hour
Administrative Support | 0 | 0 | 0 | 0 | 0
Engineer - Environmental | 0 | 0 | 0 | 0 | 0
Engineer - Water Resources | 0 | 0 | 0 | 0 | 0
Instrumentation Specialist - Senior | 0 | 0 | 0 | 0 | 0
Hydrologist - Junior Technician | 0 | 0 | 0 | 0 | 0
Hydrologist - Intermediate Technician | 0 | 0 | 0 | 0 | 0
Hydrologist - Senior | 0 | 0 | 0 | 0 | 0
Hydrometic -Senior | 0 | 0 | 0 | 0 | 0
Hydrometic - Intermediate Technician | 0 | 0 | 0 | 0 | 0
Hydrometic - Junior Technician | 0 | 0 | 0 | 0 | 0
Labourer - General | 0 | 0 | 0 | 0 | 0
Maintenance, Upgrade and Repair Costs - Materials and Supplies
Cost plus % markup | 0%
Annual Price Escalators
On the anniversary of the commencement of the commencement of the Contract, Proponents can request an annual increase based on the percentage amounts outlined below.
Description | Year Two | Year 3 | Year 4 | Year 5
Percentage | 0% | 0% | 0% | 0%
All Inclusive Cost
Subtotal without ad hoc or troubleshooting | Year 1 | Year 2 | Year 3 | Year 4 | Year 5
Site: Soames Creek – Low Flow Weir (Upstream of Marine Drive) | 0 | 0 | 0 | 0 | 0
Site: Soames Creek – Chapman Creek – Reach 3 | 0 | 0 | 0 | 0 | 0
Site: Chapman Creek – WSC / Downstream of Intake | 0 | 0 | 0 | 0 | 0
Site: Chapman Creek Upper | 0 | 0 | 0 | 0 | 0
Hydrometric Data Collection Processing Conversion | 0
Five Year Total without ad hoc or troubleshooting |  |  |  |  | 0